Development of a Low Temperature Capacitive Dilatometer for the

User Community at the NHMFL

Schmiedeshoff, G.M. (Occidental College, Physics) and Park, J.-H. (NHMFL, FSU)

Introduction

Cryogenic capacitive dilatometers [1] are usually designed to run in hard vacuum or in a low-pressure helium exchange gas because immersing the dilatometer in any medium can lead to dielectric effects which can be significant [2].  Experimental challenges also include thermal contact between the sample, thermometer, and refrigerant.  These challenges are exacerbated at the lowest temperatures, especially in the presence of high magnetic fields.

We are developing a new capacitive dilatometer to address these challenges.  The new dilatometer is designed for operation when immersed in the liquid helium of an operating dilution refrigerator.  The capacitor gap is sealed to prevent dielectric material (liquid helium) from entering.  The sample and thermometer are in direct contact with the liquid helium to maximize thermal contact and minimize thermal gradient effects.

Results and Discussion

The new design has been tested in the top-loading dilution refrigerators of the Millikelvin facility.  The capacitor gap is superfluid tight and the sensitivity is somewhat higher than expected.  However the cell-effect, the response of the “empty” cell to varying temperatures and/or magnetic fields, is larger than expected due to its sensitivity to pressure variations in the mixing chamber. Variations in design and construction materials are being explored to reduce the cell-effect.   A new system, compatible with SCM-1, has been developed which permits calibration (in a manner similar to that described in Ref. [1]) in situ, at the lowest temperatures and in the highest magnetic fields.   Preliminary data suggests that reliable operation to temperatures below 20 mK, in the highest dc fields available at the NHMFL, is an achievable goal.

Acknowledgements

A portion of this work was performed at the National High Magnetic Field Laboratory, which is supported by National Science Foundation Cooperative Agreement No. DMR-1157490, the State of Florida, and the U.S. Department of Energy.  Work at Occidental College is supported by the National Science Foundation under DMR-1408598. We gratefully acknowledge support from the UCGP and Visiting Scientist Program of the NHMFL.

References

[1]	Schmiedeshoff, G.M., et al., Rev. Sci. Instrum, 77, 123907 (2006).

[2]	Schmiedeshoff, G.M., et al., Phil. Mag., 89, 2071 (2009).