21 Tesla Fourier Transform Ion Cyclotron Resonance Mass Spectrometer:

A National Resource for Ultrahigh Resolution Mass Analysis

Hendrickson, C.L. (NHMFL; FSU, Chemistry & Biochemistry); Quinn, J.P.; Kaiser, N.K.; Smith, D.F.; Blakney, G.T. (NHMFL); Chen, T.; Marshall, A.G. (NHMFL & FSU, Chemistry & Biochemistry); Weisbrod, C.R. (NHMFL) and Beu, S.C. (S.C. Beu Consulting)

Introduction

We describe the design and initial performance of the first 21 tesla Fourier transform ion cyclotron resonance mass spectrometer. The 21 tesla magnet is the highest field superconducting magnet ever used for FT-ICR and features high spatial homogeneity, high temporal stability, and negligible liquid helium consumption.

Experimental

The instrument includes a commercial dual linear quadrupole trap front end that features high sensitivity, precise control of trapped ion number, and collisional and electron transfer dissociation. A third linear quadrupole trap offers high ion capacity and ejection efficiency, and rf quadrupole ion injection optics deliver ions to a novel dynamically harmonized ICR cell.

Results and Discussion

Mass resolving power of 150,000 (m/Δm50%) is achieved for bovine serum albumin (66 kDa) for a 0.38 second detection period (see Figure 1), and greater than 2,000,000 resolving power is achieved for a 12 second detection period. Externally calibrated broadband mass measurement accuracy is typically less than 150 ppb rms, with resolving power greater than 300,000 at m/z 400 for a 0.76 second detection period. Combined analysis of electron transfer and collisional dissociation spectra results in 68% sequence coverage for carbonic anhydrase. The instrument is part of the NSF High-Field FT-ICR User Facility and is available free of charge to qualified users.

Acknowledgments

This work was performed at the National High Magnetic Field Laboratory, which is supported by National Science Foundation Cooperative Agreement No. DMR-1157490 and the State of Florida.

Additional support was provided by the NSF Division of Chemistry (CHE-1016942 and CHE-1019193).

References

[1]	Hendrickson, C. L., et al., J. Am. Soc. Mass Spectrom. 26, 1626-1632 (2015).

Fig. 1.  Single-scan electrospray FT-ICR mass spectrum of the isolated 48+ charge state of bovine serum albumin following a 0.38 s detection period. Mass resolving power is approximately 150,000, and the signal-to-noise ratio for  the highest magnitude peak is greater than 150:1. The ion accumulation period was 250 ms and the ion number target was 5,000,000.

.