Improved Ion Optics for Introduction of Ions into a 9.4 Tesla Fourier Transform Ion Cyclotron Resonance Mass Spectrometer

Chen, Y. (NHMFL); Leach, F.E, III (Pacific Northwest National Laboratory); Kaiser, N.K. (NHMFL); Dang, X. (FSU, Chemistry & Biochemistry); Igrahim, Y.M.; Norheim, R.V.; Anderson, G.A.; Smith, R.D. (Pacific Northwest National Laboratory) and Marshall, A.G. (NHMFL & FSU, Chemistry & Biochemistry)

Introduction

Enhancements to the ion source and transfer optics of our 9.4 T FT-ICR mass spectrometer have resulted in improved ion transmission efficiency for more sensitive mass measurement of complex mixtures at the MS and MS/MS levels.

Experimental

The tube lens/skimmer has been replaced by a dual ion funnel and the following octopole by a quadrupole for reduced ion cloud radial expansion before transmission into a mass-selective quadrupole (see Figure 1).

Results and Discussion

The rate of ion accumulation is higher by an order of magnitude for the funnel/quadrupole relative to the tube lens/skimmer/octopole [1].

Acknowledgments

A portion of this work was performed at the National High Magnetic Field Laboratory, which is supported by National Science Foundation Cooperative Agreement No. DMR-1157490, and the State of Florida.

Additional support was provided by NSF CHE-1019193, and (to R.D.S.) NIH P41 GM103493. The ion funnels were designed and assembled in the W. R. Wiley Environmental Molecular Sciences Laboratory (EMSL), a DOE national scientific user facility at the Pacific Northwest National Laboratory (PNNL). PNNL is operated by Battelle for the DOE under contract DE-AC05-76RL0 1830.

References

[1]	Chen, Y., et al., J. Mass Spectrom. 50, 280-284 (2015)

Fig. 1. Schematic diagrams of the front end of the 9.4 T FT-ICR mass spectrometer. Top: The prior instrument configuration with tube lens/skimmer followed by an octopole ion guide. Bottom: Current instrument configuration with dual capillaries and PNNL-designed dual ion funnel followed by a quadrupole ion guide.