Gradient Eddy Current Pre-emphasis with Uniform Excitation of Eigenmodes

Kulam, M., Sarntinoranont, M. (UF, Mech. & Aerospace Eng); Lonser, R. (OSU, Neurosurgery); Brey, W.W. (NHMFL) and Mareci, T.H. (UF, Biochemistry)

Introduction

For pulsed field gradient (PFG) magnetic resonance (MR) experiments, characterizing the magnetic fields produced by gradient switching is important for correcting errors in the gradient moments caused by eddy currents and mechanical oscillations. Errors in gradient moments result in image distortions in k-space, and inaccuracies in q-space imaging, especially while measuring both very low diffusivities and flow velocities, which often require fast gradient switching rate. Assuming the gradient system is linear and time invariant, accurate characterization of the system requires the time derivative of a test gradient pulse mimic that of an impulse function which excite all the gradient eigenmodes uniformly. We introduce a new test pulse, called the Fresnel pulse shown in the next section, whose derivative is a chirp function, and has a more uniform spectrum over a prescribed bandwidth, like an impulse function.

Experimental

MR measurements were performed in the Advanced Magnetic Resonance Imaging and Spectroscopy Facility using an Oxford 4.7T magnet and Agilent VNMRS console, running VnmrJ 3.1A software, with shielded gradients (Model BFG-200/115-S14, Resonance Research, Inc. Billercia, MA) and a 38 mm ID quadrature birdcage transmit/receive coil. A 3D fast gradient echo imaging sequence was modified by adding a test gradient pulse with a variable delay time before RF excitation. Zeroing the gradient amplifier pre-emphasis unit and using a 28 mm DSV sample filled with distilled deionized water, the sequence was used to characterize the X-gradient with the following imaging parameters; TR/TE: 450/0.3 ms, flip angle: 16o, matrix size: 32 x 16 x 16, field of view: 32 mm x 32 mm x 32 mm, Fresnel gradient peak amplitude/ duration/ bandwidth: 500 mT/m /4 ms /2 KHz, and delay times ranged from 0 to 4 ms in steps of 0.1 ms. For a given delay time, two measurements were acquired with opposite test gradient polarities to account for non-eddy current related phase changes.

A 3D time varying eddy current magnetic field map was obtained from the phase resulting from the complex division of the two-opposite gradient-polarity images acquired at each delay time.  The resulting magnetic field was fit to a trilinear equation in space to account for the constant and gradient fields generated after the test pulse whose strengths were then fit to a decaying oscillatory multi-exponential model in time to obtain the eddy current amplitudes and time constants, and mechanical resonance frequencies.

Results and Discussion

The modeled time-dependence of the Z-gradient field generated after the X-axis Fresnel pulse is shown below. The results showed eddy currents mainly on the pulsed X-axis with a time constant of 0.5 ms, and low amplitude mechanical oscillations at roughly 500 Hz, clearly seen in the static field direction (B0) and z-axis.

Conclusions

We have presented a method to characterize gradient switching by uniform bandlimited excitation of eigenmodes using a Fresnel-shaped gradient pulse. It could be an important step in reducing eddy current artifacts in PFG experiments for the measurement of very slow diffusion and flow. Applying optimal pre-emphasis to the gradient amplifier control, measurement of the remaining residual eddy current response will allow further improvements through the design of the pulsed gradient waveform to overcome these residual eddy currents.

Acknowledgements

A portion of this work was performed at the AMRIS Facility of the National High Magnetic Field Laboratory, which is supported by National Science Foundation Cooperative Agreement No. DMR-1157490 and the State of Florida.  In addition, this work is supported in part by the NIH/NCATS Clinical and Translational Science Award UL1 TR000064.